Manchester Institute of Biotechnology  - Risk Assessment

| Date: 14/3/2016 | Assessed by: Thomas Fallows | Validated by: | Location: MIB | Review date: 14/3/2017 |
| --- | --- | --- | --- | --- |

Task: Induce heat in samples by means of a magnetic field generator.

| Activity | Hazard | Person(s) in danger | Existing measures to control risk | Risk rating | Result |
| --- | --- | --- | --- | --- | --- |
| Magnetic field generator | Electrical hazard – causing electric shock and electrical burns | Operator and any in close proximity | Ensure all electrical equipment is PAT complaint. The equipment must be visually inspected before use. Any faults/warning light indications must be reported to Simon Webb, and the equipment must be taken out of use (and labelled) until it is repaired. The equipment must always be used in accordance to the manufacturers instructions. All users must be fully trained on the safe working of the equipment, and must be supervised until deemed competent. | L | A |
| Sample handling and magnetic field generator | Hot sample and equipment – burns to hands. |  | The following items of PPE must be worn: Howie-style laboratory coat and BS EN166 compliant eye protection (chemical splash proof safety glasses). A selection of safety glasses and goggles are available from MIB Stores; users are advised to visit Stores and select eye protection which fits well and is comfortable to use. Regular lab inspections monitor the wearing of PPE; users found not to be wearing PPE when the risk assessment states that it must be worn will be subject to the MIB compliance policy. Samples are normally placed within a small glass beaker, then placed inside the AMF coil. This is not always possible, and it is acceptable to place the sample directly within the coil. Samples must be allowed to cool before handling; tongs and thermal gloves are provided for use if necessary. | L | A |
| Sample handling | Chemical hazard |  | The majority of samples used in the AMF are non hazardous (eg. Inert polymers). However, all samples are assessed before use. Any which are hazardous will be COSHH risk assessed before use, the assessment read and signed, and the control measures followed. Some samples are provided by users from other Schools, who access our equipment. Often, these samples are mounted on aluminium foil. It is essential that all aluminium is removed before the sample is placed within the AMF, as the magnetic field would cause the temperature of the aluminium to increase and could ignite it. | L | A |

Authorisation by PI I confirm that I have considered and understand the experiment and the associated hazards. I am satisfied that all of the hazards have been identified and that the control measures to be followed will reduce the risks to acceptable levels. Print name: Signed: Date:

Declaration by researcher

I confirm that I have read this Risk Assessment and that I understand the hazards and risks involved and will follow all of the safety procedures stated.

Declaration by PI

I confirm that the researcher who has signed below is competent to undertake the work. My counter-signature indicates that I am happy for the work to proceed.

| Name (please print) | signed | PI countersignature | date |
| --- | --- | --- | --- |